Malaysian Journal of Analytical Sciences Vol 24 No 6 (2020): 927 - 939

OPTIMISATION AND VALIDATION OF ULTRAHIGH PERFORMANCE LIQUID CHROMATOGRAPHY METHOD FOR QUANTIFICATION OF

25-HYDROXYVITAMIN D IN MATERNAL PLASMA

(Pengoptimuman dan Pengesahan Kaedah Kromatografi Cecair Berprestasi Ultra Tinggi untuk Kuantifikasi 25-hidroksivitamin dalam Plasma Ibu Mengandung)

Siew Siew Lee and Su Peng Loh*

Department of Nutrition,

Universiti Putra Malaysia, 43400 UPM Serdang, Selangor, Malaysia

*Corresponding author:  sploh@upm.edu.my

Received: 19 August 2020; Accepted: 25 October 2020; Published:  xx December 2020

Abstract

This study aimed to optimise and validate a simple and efficient sample preparation and extraction method for the quantification of 25-hydroxyvitamin D (25OHD) in the maternal plasma sample. Different sample preparation methods and precipitation reagent-to sample ratio used by previous studies were compared. An ultrahigh performance liquid chromatography method was developed and validated for simultaneous quantification of 25OHD2 and 25OHD3. The chromatographic separation was achieved using the COSMOCORE 2.6Cholester Column and methanol and 0.1% formic acid (79:21, %v/v) as mobile phase at a flow rate of 0.3 mL/min and diode array detection at 264 nm. The results demonstrated that a recovery of approximately 100% could be achieved by extracting samples using 1 mL of hexane, vortex for 10s and a total number of three extraction steps. The precipitation reagent-to-sample ratio of 2.8 was optimum for the quantification of 25OHD2 and 25OHD3 in a pooled maternal sample. The value obtained for validation parameters meets the criteria of the Recommendations and Acceptance Criteria for Bioanalytical Method Validation by the Food and Drug Administration. The results showed that this method could be applied for routine quantification of 25OHD, particularly in the maternal plasma sample.

Keywords:  25-hydroxyvitamin D, maternal, ultrahigh performance liquid chromatography, extraction

Abstrak

Kajian ini bertujuan mengoptimumkan dan mengesahkan kaedah penyediaan dan pengekstrakan yang mudah dan cekap untuk pengkuantitian 25-hidroksivitamin D (25OHD) di dalam sampel plasma ibu mengandung. Kaedah kromatografi cecair berprestasi ultra tinggi telah dibangunkan dan disahkan untuk pengkuantitian 25OHD2 and 25OHD3 secara serentak. Pemisahan kromatografi telah dicapai melalui turus COSMOCORE 2.6Cholester dan metanol: 0.1% asid formik (79:21, %v/v) sebagai fasa bergerak pada kadar aliran 0.3 mL/min dan pengesahan susunan diod pada 264 nm. Keputusan kajian menunjukkan bahawa kaedah pengekstrakan terbaik yang memberi kadar pemulihan yang hampir dengan 100% adalah mengekstrak sampel dengan menggunakan 1mL heksana, vorteks selama 10 s dan sejumlah 3 kali bagi langkah pengekstrakan. Kadar reagen pemendakan kepada sampel sebanyak 2.8 adalah optimum bagi pengekstrakan 25OHD2 dan 25OHD3 dalam sampel plasma ibu mengandung. Nilai yang diperolehi untuk parameter validasi adalah memenuhi kriteria cadangan dan kriteria penerimaan untuk pengesahan kaedah bioanalitikal oleh pentadbiran makanan dan ubat-ubatan (FDA). Hasil kajian menunjukkan bahawa kaedah yang diperolehi boleh diguna untuk rutin pengkuantitian 25OHD dalam sampel plasma, terutamanya sampel plasma dari ibu mengandung.

Kata kunci:  25-hidroksivitamin D, ibu mengandung, kromatografi cecair berprestasi ultra tinggi, pengekstrakan

References

- Institute of Medicine. (2011). Dietary Reference Intakes for Calcium and Vitamin D. The National Academies Press, Washington, DC: pp. 96.

- Scientific Advisory Committee on Nutrition (2016). SACN vitamin D and Health report. https://www.gov.uk/government/publications/sacn-vitamin-d-and-health-report. [Access online 20 January 2019].

- Wallace, A. M., Gibson, S., de la Hunty, A., Lamberg-Allardt, C. and Ashwell, M. (2010). Measurement of 25-hydroxyvitamin D in the clinical laboratory: Current procedures, performance characteristics and limitations. Steroids, 75(7): 477-488.

- Bikle, D., Gee, E., Halloran, B., Kowalski, M. A., Ryzen, E. and Haddad, J. G. (1986). Assessment of the free fraction of 25-hydroxyvitamin D in serum and its regulation by albumin and the vitamin D-binding protein. The Journal of Clinical Endocrinology & Metabolism, 63(4): 954-959.

- Le Goff, C., Cavalier, E., Souberbielle, J. C., Gonzalez-Antuna, A. and Delvin, E. (2015). Measurement of circulating 25-hydroxyvitamin D: A historical review. Practical Laboratory Medicine, 2: 1-14.

- Bikle, D. D., Gee, E., Halloran, B. and Haddad, J. G. (1984). Free 1,25-dihydroxyvitamin D levels in serum from normal subjects, pregnant subjects, and subjects with liver disease. The Journal of Clinical Investigation, 74(6): 1966-1971.

- van Hoof, H. J., de Sevaux, R. G. and van Baelen, H. (2001). Relationship between free and total 1,25-dihydroxyvitamin D in conditions of modified binding. European Journal of Endocrinology, 144(4): 391 – 396.

- Jones, K. S., Assar, S., Prentice, A. and Schoenmakers, I. (2016). Vitamin D expenditure is not altered in pregnancy and lactation despite changes in vitamin D metabolite concentrations. Scientific Reports, 6, 26795.

- Freeman, J., Wilson, K., Spears, R., Shalhoub, V. and Sibley, P. (2014). Influence of vitamin D binding protein on accuracy of 25-hydroxyvitamin D measurement using the ADVIA centaur Vitamin D total assay. International Journal of Endocrinology, 2014: 1-12.

- Heijboer, A. C., Blankenstein, M. A., Kema, I. P. and Buijs, M. M. (2012). Accuracy of 6 routine 25-hydroxyvitamin D assays: Influence of Vitamin D binding protein concentration. Clinical Chemistry, 58(3): 543-548.

- Cavalier, E., Wallace, A. M., Knox, S., Mistretta, V. I., Cormier, C. and Souberbielle, J. C. (2008). Serum vitamin D measurement may not reflect what you give to your patients. Journal of Bone and Mineral Research, 23(11): 1864-1865.

- Lankes, U., Elder, P. A., Lewis, J. G. and George, P. (2015). Differential extraction of endogenous and exogenous 25-OH-vitamin D from serum makes the accurate quantification in liquid chromatography-tandem mass spectrometry assays challenging. Annals of Clinical Biochemistry, 52(Pt 1): 151-160.

- Centers for Disease Control and Prevention (2010). Laboratory procedure manual; 25-hydroxyvitamin D3, 3-epi-25-hydroxyvitamin D3, 25-hydroxyvitamin D2. Access from https://wwwn.cdc.gov/Nchs/Data/Nhanes/2009-2010/LabMethods/VID_F_met_Vitamin_D.pdf. [Access 30 September 2018]

- Turpeinen, U., Hohenthal, U. and Stenman, U. H. (2003). Determination of 25-hydroxyvitamin D in Serum by HPLC and Immunoassay. Clinical chemistry, 49(9): 1521 – 1524.

- Chin, S. F., Osman, J. and Jamal, R. (2018). Simultaneous determination of 25-hydroxyvitamin D2 and 25-hydroxyvitamin D3 in human serum by ultra-performance liquid chromatography: An economical and validated method with bovine serum albumin. Clinica Chimica Acta, 485: 60-66.

- US Department of Health Human Services Food and Drug Administration, Center for Drug Evaluation and Research, Center for Veterinary Medicine. (2001). Guidance for Industry: Bioanalytical method validation (Revised May 2018).

- International Conference on Harmonization (ICH) of Technical Requirement for the Registration of Pharmaceuticals for Human Use. ICH Harmonised Tripartite Guideline: Validation of analytical procedures: text and methodology Q2 (R1).

- Kand'ar, R. and Zakova, P. (2009). Determination of 25-hydroxyvitamin D3 in human plasma using HPLC with UV detection based on SPE sample preparation. Journal of Separation Science, 32(17): 2953-2957.

- Chen, H., McCoy, L. F., Schleicher, R. L. and Pfeiffer, C. M. (2008). Measurement of 25-hydroxyvitamin D3 (25OHD3) and 25-hydroxyvitamin D2 (25OHD2) in human serum using liquid chromatography-tandem mass spectrometry and its comparison to a radioimmunoassay method. Clinica Chimica Acta, 391(1-2): 6-12.

- Abu el Maaty, M. A., Hanafi, R. S., Aboul-Enein, H. Y. and Gad, M. Z. (2015). Design-of-experiment Approach for HPLC Analysis of 25-hydroxyvitamin D: A comparative assay with ELISA. Journal of Chromatographic Science, 53(1): 66-72.

- Mineva, E. M., Schleicher, R. L., Chaudhary-Webb, M., Maw, K. L., Botelho, J. C., Vesper, H. W. and Pfeiffer, C. M.  (2015). A candidate reference measurement procedure for quantifying serum concentrations of 25-hydroxyvitamin D(3) and 25-hydroxyvitamin D(2) using isotope-dilution liquid chromatography-tandem mass spectrometry. Analytical and Bioanalytical Chemistry, 407(19): 5615- 5624.

- Stepman, H. C., Vanderroost, A., Van Uytfanghe, K. and Thienpont, L. M. (2011). Candidate reference measurement procedures for serum 25-hydroxyvitamin D3 and 25-hydroxyvitamin D2 by using isotope-dilution liquid chromatography-tandem mass spectrometry. Clinical Chemistry, 57(3): 441-448.

- Hymoller, L. and Jensen, S. K. (2011). Vitamin D analysis in plasma by high performance liquid chromatography (HPLC) with C(30) reversed phase column and UV detection-easy and acetonitrile-free. Journal of Chromatography. A, 1218(14): 1835-1841.

- Nurmi, T., Tuomainen, T. P., Virtanen, J., Mursu, J. and Voutilainen, S. (2013). High-performance liquid chromatography and coulometric electrode array detector in serum 25-hydroxyvitamin D(3) and 25-hydroxyvitamin d(2) analyses. Analytical Biochemistry, 435(1): 1-9.

- Maunsell, Z., Wright, D. J. and Rainbow, S. J. (2005). Routine isotope-dilution liquid chromatography-tandem mass spectrometry assay for simultaneous measurement of the 25-hydroxy metabolites of vitamins D2 and D3. Clinical Chemistry, 51(9): 1683-1690.

- Franke, A. A., Morrison, C. M., Custer, L. J., Li, X. and Lai, J. F. (2013). Simultaneous analysis of circulating 25-hydroxy-vitamin D3, 25-hydroxy-vitamin D2, retinol, tocopherols, carotenoids, and oxidized and reduced coenzyme Q10 by high performance liquid chromatography with photo diode-array detection using C18 and C30 columns alone or in combination. Journal of Chromatography. A, 1301: 1-9.

- Hrvolová, B., Martínez-Huélamo, M., Colmán-Martínez, M., Hurtado-Barroso, S., Lamuela-Raventós, R. M. and Kalina, J. (2016). Development of an advanced HPLC-MS/MS method for the determination of carotenoids and fat-soluble vitamins in human plasma. International Journal of Molecular Sciences, 17(10): 1719.

- Lensmeyer, G. L., Wiebe, D. A., Binkley, N. and Drezner, M. K. (2006). HPLC method for 25-hydroxyvitamin D measurement: Comparison with contemporary assays. Clinical Chemistry, 52(6): 1120 – 1126.

- Garg, U., Munar, A., Frazee, C. nd Scott, D. (2012). A simple, rapid atmospheric pressure chemical ionization liquid chromatography tandem mass spectrometry method for the determination of 25-hydroxyvitamin D2 and D3. Journal of Clinical Laboratory Analysis, 26(5): 349-357.

- Bruce, S. J., Rochat, B., Beguin, A. Pesse, B., Guessous, I., Boulat, O. and Henry, H. (2013). Analysis and quantification of vitamin D metabolites in serum by ultra-performance liquid chromatography coupled to tandem mass spectrometry and high-resolution mass spectrometry-a method comparison and validation. Rapid Communications in Mass Spectrometry, 27(1): 200-206.

- Eyles, D., Anderson, C., Ko, P. Jones, A., Thomas, A., Burne, T. and McGrath, J. (2009). A sensitive LC/MS/MS assay of 25OH vitamin D3 and 25OH vitamin D2 in dried blood spots. Clinica Chimica Acta, 403(1-2): 145-151.

- Zhang, S. W., Jian, W., Sullivan, S. Sankaran, B., Edom, R. W., Weng, N. and Sharkey, D. (2014). Development and validation of an LC-MS/MS based method for quantification of 25 hydroxyvitamin D2 and 25 hydroxyvitamin D3 in human serum and plasma. Journal of Chromatography. B, Analytical Technologies in the Biomedical and Life Sciences, 961: 62-70.

- Karras, S. N., Kotsa, K., Angeloudi, E., Zebekakis, P. and Naughton, D. P. (2017). The Road Not So Travelled: Should Measurement of Vitamin D Epimers During Pregnancy Affect Our Clinical Decisions? Nutrients, 9(2): 90.

- Aghajafari, F., Field, C. J., Rabi, D., Kaplan, B. J., Maggiore, J. A., O’Bierne, M., Hanley, D. A., Eliasziw, M., Dewey, D., Ross, S. and Apron T. (2016). Plasma 3-epi-25-hydroxycholecalciferol can alter the assessment of vitamin d status using the current reference ranges for pregnant women and their newborns. The Journal of Nutrition, 146(1): 70-75.

- Yazdanpanah, M., Bailey, D., Walsh, W., Wan, B. and Adeli, K. (2013). Analytical measurement of serum 25-OH-vitamin D(3), 25-OH-vitamin D(2) and their C3-epimers by LC-MS/MS in infant and pediatric specimens. Clinical Biochemistry, 46(13-14): 1264-1271.

- Kiely, M., O'Donovan, S. M., Kenny, L. C., Hourihane, J. O., Irvine, A. D. and Murray, D. M. (2017). Vitamin D metabolite concentrations in umbilical cord blood serum and associations with clinical characteristics in a large prospective mother-infant cohort in Ireland. Journal of Steroid Biochemistry and Molecular Biology, 167: 162-168.