210PO DETERMINATION IN URINE SAMPLES AMONG RADIATION WORKERS BY ALPHA SPECTRUM ANALYSIS

(Penentuan 210Po dalam Sampel Urin di Kalangan Pekerja Sinaran

Melalui Analisis Spektrum Alpha)

Nurrul Assyikeen Md. Jaffary*, Nurul Zakila Azlan, Jalal Sharib @ Sarip, Mohd Tarmizi Ishak,

Zulkifli Daud, Mohd Izwan Abdul Adziz

Radiochemistry and Environmental Laboratory,

Waste and Environmental Technology Division,

Malaysian Nuclear Agency,

Bangi, 43000 Kajang, Malaysia

+Corresponding author: assyikeen@nuclearmalaysia.gov.my

Abstract

The radioactivity of 210Po in the urine sample of radiation workers working in radiochemical laboratory, Malaysian Nuclear Agency, was determined. Collected urine samples have to go through the hot acid dissolution process followed by the auto-deposition technique using HCl as the solution. Measurements using the alpha spectrometry system showed good 209Po retrieval, while 210Po determination has no interference by other natural radionuclides. The specific radioactivity of 210Po was determined from 1.33 mBqL-1 to 49.84 mBqL-1. This preliminary study was conducted to obtain reference data as part of methods development for further investigations in emergencies. The assessment of the data findings in terms of radiation safety was done. Such radioactivity measurement procedure would also be applying for the routine 210Po analysis in the urine.

Keywords: polonium isotopes, radioactivity, radiochemical separation, alpha spectrometry, bioassay

Abstrak

Radioaktiviti 210Po dalam sampel urin pekerja radiasi yang bekerja di makmal radiokimia, Agensi Nuklear Malaysia telah ditentukan. Sampel urin yang dikumpulkan mesti melalui proses pembubaran asid panas diikuti dengan teknik pemendapan-auto menggunakan HCl sebagai medium larutan. Pengukuran menggunakan sistem spektrometri alpha menunjukkan dapatan semula 209Po yang baik, sementara penentuan 210Po tiada gangguan daripada kehadiran radionuklida semula jadi yang lain. Radioaktiviti spesifik 210Po telah ditentukan dari 1.33 mBqL-1 hingga 49.84 mBqL-1. Kajian awal ini dilaksanakan untuk mendapatkan data rujukan sebagai sebahagian dari pengembangan kaedah bagi situasi kecemasan. Penilaian ke atas dapatan data dari segi keselamatan radiasi telah dilakukan. Prosedur pengukuran radioaktif seperti ini juga digunakan untuk aktiviti analisis rutin 210Po dalam sampel urin.

Kata kunci: isotop polonium, radioaktiviti, pemisahan radiokimia, spektrometri alfa, bioasai

# Introduction

210Pb and 210Po are associated with the 238U decay chain, a naturally occurring radionuclide. 210Pb and its daughter 210Bi are beta sources with a half-life of 22.3 years and 5.0 days, respectively. Meanwhile, 210Po, the granddaughter of 210Pb, is an alpha transmitter. The half-life of 210Po is only 138.4 days, and it is the lengthiest polonium radioisotope. 210Po occurrences have been continually supplied in the environment by 210Pb, which is technologically enhanced from natural radioactive materials (TENORM) [1].

Owing to the natural background radiation, 210Po contributes significantly to human internal dosage when absorbed into the body by inhalation or ingestion and is therefore  essential to be monitored [1]. An excessive dose of 210Po may cause acute radiation syndrome (ARS), which includes nausea, vomiting, diarrhea in the prodromal period, and later symptoms, including hair loss and bleeding.  A lower dose can cause long-term organ failure, which could lead to death. The medical sequence can be defined  when intake of 210Po more than 1 GBq (109 Bq) occurs [2].

Historically, 210Po has also been linked to the death of Mr. Alexande Litvinenko, a former Soviet spy. He died on 23rd November 2006, 22 days after consuming a cup of tea laced with 210Po. The assassins left radioactive traces all over London. Following this attack, London Health officials used alpha spectrometry to track hundreds of urine samples from internally infected individuals [2, 3]. Despite the fact that 210Po can be easily determined by measuring alpha radiation emissions, there is a shortage of reference on 210Po in human tissues and body fluids used as a baseline in determining whether or not the individual would be contaminated [4].

The bioassay technique was used to measure the amount of actinide present in the body based on the rate of excretion of body fluids. In an emergency of high 210Po activity spread, urine samples of individuals dealing specifically with 210Po from open sources should be analysed. It is critical to act rapidly in an emergency to identify radiological threats, particularly to the general public. Urine bioassays are often used to assess internal 210Po exposure [2], which can be separated into two types; spontaneous plating [5] and micro-precipitation technique [6]. However, the pre-concentration approach is tedious and time-consuming, which may not be appropriate in an emergency. Therefore, to measure 210Po in urine samples, new and quick analysis methods are needed.

In this study, sample selection from the radiochemical laboratory of the Malaysian Nuclear Agency was used as a reference group to determine 210Po. This paper defines the procedure used for determining 210Po in urine samples, which involved wet decomposition followed by automated deposition for alpha spectrometry. Furthermore, this research focuses on the factors that affect polonium determination and the need to improve the method. Finally, the study results of 210Po value will be used as the primary baseline reference data, and assessment in terms of radiation safety were done.

Materials and Methods

## Materials

209Po was purchased from Perkin Elmer and was diluted to 30Bq. The diluted concentration was calibrated by preparing a set of a known amount of standards and using a calibration curve to create a radioactivity relationship between the amount spiked and an instrument response. Nitric acid (65% purity) and hydrogen peroxide (30% purity) were purchased from MERCK, used for acid dissolution. Hydrochloric acid (37% purity) was supplied by Scharlau, and distilled water was used for acid dilution in preparing final solution of 2M HCl for auto-deposition process. Mixed standard alpha reference sources (238U, 234U, 239Pu and 241Am) were obtained from ANALYTICS used to calibrate the alpha spectrum energy for the alpha spectrometry system.

## Sampling and sample analysis

A 24-hour urine specimen obtained from 6 radiochemical workers were used to examine. 50 mL of the obtained urine sample, moved to a 100 mL glass beaker and diluted with 50 mL of distilled water. Each sample was made in triplicate and heated to just below 90 ºC until almost dry. The known amount (approximately 6 Bq) of 209Po as the tracer was added. In the urine bioassay method, 210Po has often been corrected using a 208Po tracer or 209Po to assess the recovery due to any depletion or cross contamination of 210Po during the sample analysis. In this study, 209Po was used as a tracer due to the advantage in energy isolation from the 210Po peak. The 208Po is not widely used as a tracer since its peak energy is closed to 210Po and can be interference in mistake (Table 1).

Table 1. 208Po, 209Po, and 210Po decay data

| Radionuclides | Half-life (t1/2) | Disintegration Modes | Eα (MeV) | Intensity (%) |
| --- | --- | --- | --- | --- |
| 208Po | 2.898y | α:99.99777% | 5.1149 4.220 | 99.9956 4.220 |
| 209Po | 102y | α:99.52% | 4.885 4.883 4.622 4.310 4.110 | 19.7 79 0.551 1.5x10-4 5.6x10-4 |
| 210Po | 138.376 | α:100 | 5.40746 4.60436 | 99.99876 1.24x10-3 |

References: [7]

Complete dissolution was carried out with 10 mL of 65% nitric acid and 1 mL of 30% hydrogen peroxide until all the organics had been decomposed. The solution was evaporated until almost dry after dissolution, and the digestive residue was converted to chloride form by applying 10 mL 37% hydrochloric acid. Once the evaporation process was completed, 3.5 mL of 37% hydrochloric acid was applied directly to dissolve the residue. The solution was then rendered to 0.5M hydrochloric acid medium by adding 80 mL of distilled water.

### Auto plating

In each sample, hydroxylamine chloride and Bi carrier were added, which were then gently heated on a hot plate and dissolved before being placed on the silver disk as seen in Figure 1. 210Pb and its daughter 210Bi must be entirely separated to maintain the secular equilibrium between 210Bi and 210Po with no interference of 210Pb [3]. 4 hours of plating is needed to achieve optimum recovery [8]. Reducing the plating time (at room temperature) to speed up sample processing throughput could be an alternative, but it will reduce the recovery and increase the MDA. In this case, the volume of tracer used must be adjusted to account for the reduction of recovery.

Figure 1.  Image of an auto-deposition setup

### Active source measurement using alpha spectrometry system

The 450 mm2 silver discs active surface was then measured using an alpha spectrometry system consisting of Alpha Passivated Implanted Planar Silicon (PIPS) detectors, as shown in Figure 2(a). Compared to other radiometric methods, such as γ-ray spectrometry, alpha spectrometry has a distinct advantage: the background is very low. The efficiency of more than 20% was measured based on efficiency measurement of ANALYTICS mixed standard alpha reference sources for a source to the sensor distance less than 1 cm (shelve number 1 in the alpha detector chamber), as shown in Figure 2(b). The detector efficiency does not vary with energy (4 MeV to 9 MeV) due to the small distance of α-peak energy. Because of this efficiency’s freedom from energy (i.e. the intrinsic efficiency of one), the precise non-standard alpha activity ratio for peak energy’s efficiency calibration may be determined [9].

Figure 2.  a) Active source placed paralleled to the detector b) Shelve placed at the same calibrated mixed alpha standard ANALYTICS source

The duration of the active source measurements in this study was 24 hours. Due to its volatility, the polonium atoms readily exit the source disc in the vacuum counting chamber, allowing a significant fraction drift to the detector. Because of the drawback of the long half-life of 209Po, extra care was taken on sample handling to prevent detector contamination. Figure 3. shows the overall 210Po analysis in the urine by wet decomposition for alpha analysis.

Auto-deposition on a silver plate at 90oC for 4hours

Sample

Measurement of 210Po by alpha-spectrometry (24-hour counting time)

Sample preparation

Source preparation

Measurement

Sample was diluted to 100ml

Spike 209Po tracer

Sample was dried up

Mixture was dissolved with 3.5 ml conc HCl

Mixture was rinsed and marked up with distilled water until it reached 80 ml solution to be 0.5M

1 ml stable Bi3+ carrier was added

1 g of hydroxylammoniumchloride was added

210Po deposition

- A sample was collected from 6 radiochemistry workers within a 24-hour urine specimen.

- 50 ml was used, each sample prepared in triplicate

Sampling

Sample digestion

Sample was added to 10 ml HNO3+1 ml H2O2

Figure 3.  A flow chart for determination of 210Po in urine samples using alpha spectrometry system

Figure 4.  Polonium alpha, 210Po spectrum with 209Po yield tracer from this study

The polonium alpha spectrum with 209Po viewed in Alpha Vision software is shown in Figure 4. A particular amount of 209Po was added to the sample to be used as an internal isotope tracer, making it possible to calculate the results of the radiochemical analysis. The result was calculated from the peak area of 209Po using equation (1), while 210Po activity was calculated from equation (2).

ηchem=Area209PoA209Po∙tm209Po.εdet								   (1)

A210Po=Area210Potηchem∙.εdet∙ms									   (2)

Where, A210Po is activity of 210Po in the sample, Bqkg-1, A209Po is activity of 209Po in the sample, Bqkg-1, Area210Po is peak area of 210Po, counts, Area209Po is peak area of 209Po, counts, t is measurement time, s, ηchem is yield of 209Po tracer,%, εdet is counting efficiency, %, m209Po is tracer mass, g, and ms is sample mass, kg.

# Results and Discussion

Urine samples

Table 2 shows the 210Po concentrations in the urine samples relative to blank samples. The average specific radioactivity measured is 17.93 mBqL-1, with an estimated relative uncertainty of 0.80 mBqL-1. The recovery data is quite good and comparable to other research [6,8]. The varies of 210Po of each individual is greatly dependent upon the daily excretion rates between individuals. Individual variations in urine volume provided to the laboratory for examination may imply that specific urine samples can lead to a collection period of less than or more than 24 hours.

Table 2.  Results of determination 209Po in urine sample

| Sample | 210Po Determination (mBqL-1±1SD) | Recovery (%) |
| --- | --- | --- |
| Blank | 2.85 | 77.59 |
| P 1 | 12.67±0.44 | 72.78 |
| P 2 | 38.42±0.26 | 74.93 |
| P 3 | 41.46±0.51 | 79.08 |
| P 4 | 12.48±0.18 | 76.32 |
| P 5 | 10.78±0.23 | 80.61 |
| P 6 | 17.86±0.57 | 78.48 |

Table 3.  Comparison of 210Po in urine with the previous study measured using alpha spectrometry system

| Concentration (mBqL-1) | Sample | References |
| --- | --- | --- |
| Average: 23.27 Min: 1.33 Max: 49.84 | Radiochemistry laboratory worker | This study |
| Min: 0.5 Max: 4.8 | Reference group of Portuguese citizens | [4] |
| 5.9 8.9 8.1 | Saudi Arabia Volunteer public - Non-smokers - Smokers - Shisha smokers | [10] |
| Min: 0.6 Max: 69.6 | 550 measurements from the general public were examined | [11] |

The results of the 210Po bioassay are compared to those of the preceding sample in Table 3. In comparison to previous investigations, the data has showed that the specific radioactivity is slightly greater than other reported research [4, 10]. The variation of 210Po in urine is influenced by the targeted person, with 210Po in urine acquired from radiation workers in this study being more susceptible to internal exposure. Radiation workers who work with TENORM should wear personal protective equipment (PPE) such as a respiratory mask to prevent inhalation or ingestion of the radioactive gas or dust to avoid serious health effects due to external and or internal exposure on workers [12]. Considering the 238U decay series from TENORM activity, workers may internally expose to the radioactive 226Ra gas by inhalation if without applying radiation protection principles. The 226Ra then decays to its granddaughter of 210Po in particulate form and is trapped and accumulated inside the body. The alpha particle will ionize the individual’s cell and cause damage to the cell and tissue. The considering pathway may give an acute health risk to the radiation worker. The reported data in this study, however, is still within the background levels of 210Po in urine which has been examined and assessed for almost 550 measurements of the general public [11].

Assessment on the detection level

The 210Po concentration in urine samples from the current study was computed into the daily excretion rate determined in mBq.d–1, as shown in Table 4. Based on the UK Health Protection Agency’s experience in assessing 210Po contamination of members of the public in London areas, the maximum committed effective dose to be considered is below 1 mSv.y–1, which is computed by 210Po daily excretion rate determined at or below 30       mBqd–1 [4]. It is well known that the 210Po levels in urine differed from person to person based on environmental and lifestyle variables [6]. Studies have indicated that more than 80% of the contribution from 210Po to the total annual background dose is derived from 210Po associated with seafood consumption [14, 15]. The accumulation of 210Po in various environmental products, including fish and shrimp, is well-reported [16, 17]. Hence, food intake across the district can also affect the data, while uptake through inhalation of air and cigarette smoke gives smaller contributions [18].

Table 4. 210Po daily excretion rate compared with other studies measured using alpha spectrometry system

| Country | Sample /Country | Average Daily Excretion Rate (mBqd-1) | References |  |
| --- | --- | --- | --- | --- |
| Malaysia | Radiochemistry laboratory worker | Average:1.16 Min: 0.07 Max: 2.49 |  | This study |
| Brazil | Monazite plant workers - Non-smokers - Smokers | 5.2 9.9 |  | [13] |
| Italy | Volunteer public - Non-smokers - Smokers | 11.8 12 |  | [5] |
| Portugal | Portuguese citizen potentially exposed to 210Po in London | 3.06 |  | [4] |
| United Kingdom | UK residents acquired internal contamination | 30 mBq in 24 hours (when computed to committed effective dose is below 1mSvyr-1) |  | [2] |

Table 5.  Level of exposure according to category

| Category | Level of Exposure | Committed Effective Dose | References |
| --- | --- | --- | --- |
| 1 | Public | 0-1 mSvyr-1 | [6] |
| 2 | Occupational | 1-20 mSv yr-1 |  |
| 3 | Emergency | >20mSv |  |

According to the category of public, occupational, and emergency, the level of exposure is tabulated in Table 5. According to the findings of the current report, the following screening thresholds were deemed to be of no significance in terms of radiation safety. The required detection level (RDL) for assessing background levels in human urine and chronic exposures is 3mBqd-1 for the general public, indicating that the Malaysian Nuclear Agency's Radiochemistry Laboratory has been far from the criteria for the assessment of 210Po in urine. Based on this study, we would like to emphasise that the suitability of experimental procedures can be ascertained as 210Po measurements with alpha spectrometry are given in a standardised manner and have been extensively used for various biological samples such as marine samples, body fluids like urine and blood [4, 19].

# Conclusions

The proposed method provided an easy way to determine 210Po concentration in a urine sample. The most common method for measuring 210Po in urine is sample digestion, Po isolation, and auto deposition, accompanied by α spectrometric counting. Urine samples are conventionally digested and evaporated in an open beaker on a hotplate in the presence of oxidising agents (acids and hydrogen peroxide).  Since radioactive compounds like 210Po are present in nature, they can also be found in the human body. To assess any contaminants, a clearer understanding of the population's baseline level is needed. The 210Po excretion rate can be used to estimate doses resulting from 210Po or a combination of 238U chain radionuclides that comprises 210Po. The 210Po of radiochemistry laboratory’s worker in this study were measured between 1.33 mBqL-1 to 49.84 mBqL-1. Finally, further research should be conducted to increase the precision and accuracy of 210Po determination on urine samples significantly. The RAS laboratory must always get ready with 210Po and other radionuclides on hand for bioassay research. This method should allow to measure low levels of 210Po and a fast turnaround for screening. In an emergency, this preliminary study was suggested for further review, considering continuous method development.

Acknowledgments

The authors thank the management of the Waste and Environmental Technology Division for supporting the implementation of this project, and the staff of Alpha Laboratory, Radiochemistry and Environmental Laboratory for their assistance in this work.

# References

- Vreček, P., Benedik, L. and Pihlar, B. (2004). Determination of 210Pb and 210Po in sediment and soil leachates and in biological materials using a Sr-resin column and evaluation of column reuse. Applied Radiation and Isotopes, 60(5): 717-723.

- Maguire, H., Fraser, G., Croft, J., Bailey, M., Tattersall, P., Morrey, M. and Troop, P. (2010). Assessing public health risk in the London polonium-210 incident, 2006. Public Health, 124(6): 313-318.

- Li, C., Sadi, B., Davis, K., Wyatt, H., Cornett, J., Kramer, G. and Canada, P. E. (2011). Is alpha spectrometry reliable for 210Po urine bioassay? Radiation Protection Dosimetry, 143(1): 106-108.

- Carvalho, F. P. and Oliveira, J. M. (2009). Bioassay of 210Po in human urine and internal contamination of man. Journal of Radioanalytical and Nuclear Chemistry, 280(2): 359-362.

- Meli, M. A., Desideri, D., Roselli, C. and Feduzi, L. (2009). 210Po determination in urines of people living in Central Italy. Journal of Environmental Radioactivity, 100(1): 84-88.

- Manickam, E., Sdraulig, S. and O’Brien, R. (2010). An improved and rapid radiochemical method for the determination of polonium-210 in urine. Australian Journal of Chemistry, 63(1): 38-46.

- IAEA (2009). A procedure for the determination of Po-210 in water samples by alpha spectrometry. IAEA Analytical Quality in Nuclear Applications Series No. 12.

- Guérin, N. and Dai, X. (2015). An emergency bioassay method for 210Po in urine. Applied Radiation and Isotopes, 103: 179-184.

- Aggarwal, S. K. (2016). Alpha-particle spectrometry for the determination of alpha emitting isotopes in nuclear, environmental and biological samples: Past, present and future. Analytical Methods, 8(27): 5353-5371.

- Al-Arifi, M. N., Alkarfy, K. M., Al-Suwayeh, S. A., Aleissa, K. A., Shabana, E. I., Al-Dhuwaili, A. A. and Al-Hassan, M. I. (2006). Levels of 210Po in blood, urine and hair of some Saudi smokers. Journal of Radioanalytical and Nuclear Chemistry, 269(1): 115-118.

- Hodgson, A. (2018). Natural levels of Polonium-210 in urine. Journal of Radiological Protection. 38: 393.

- Jaffary, N. A., Khoo, K. S., Mohamed, N. H., Yusof, M. A. W. and Mohd Fadzil, S. (2019). Malaysian monazite and its processing residue: chemical composition and radioactivity. Journal of Radioanalytical and Nuclear Chemistry, 322(2): 1097-1105.

- Lipsztein, J. L., Melo, D. R., Sousa, W., Dias da Cunha, K. M., Azeredo, A. M. G., Julião, L. and Santos, M. (2003). Norm workers: A challenge for internal dosimetry programmes. Radiation Protection Dosimetry, 105(1–4), 317-320.

- Alonso-Hernandez, C., Diaz-Asencio, M., Munos-Caravaca, A., Suarez-Morell, E. and Avila-Moreno, R. (2002). 137Cs and 210Po dose assessment from marine food in Cienfuegos Bay (Cuba). Journal of Environmental Radioactivity, 61(2): 203-211.

- Noshkin, V. E., Robison, W. L. and Wong, K. M. (1994). Concentration of 210Po and 210Pb in the diet at the Marshall Islands. Science Total of Environment, 155: 87-104.

- UNSCEAR (2000). Sources and effects of ionizing radiation, New York: pp 1.

- Alam, L. and Mohamed, C. A. R. (2011). Natural radionuclide of Po210 in the edible seafood affected by coal-fired power plant industry in Kapar coastal area of Malaysia. Environmental Health: A Global Access Science Source, 10(1): 1-10.

- Carvalho, F. P. (1995). Origins and concentrations of 222Rn, 210Pb, 210Bi and 210Po in the surface air at Lisbon, Portugal, at the Atlantic edge of the European continental landmass. Atmospheric Environment, 29(15): 1809-1819.

- García-Orellana, I. and García-León, M. (2002). An easy method to determine 210Po and 210Pb by alpha spectrometry in marine environmental samples. Applied Radiation and Isotopes, 56(4): 633-636.